## Slide 1: BSL3,ABSL3 Lab Spec. for Bidding
2020/2/23, 2020/3/24, 2020/4/14,
2020/4/19, 2020/7/7, 2020/714,
2020/7/31
PIHCMC

## Slide 2: Existing Inside Walls in No.5 Building
Remove
(inside walls and inside doors, except columns and beams)
Note) Existing columns and beams are a few and thin. So confirm structural strength carefully. And reinforce, if needed.
(refer to No.5 building drawing)

## Slide 3: Fire Resistance
Replace
Remove
Note) Windows in concrete wall should be replaced.
(refer to No.5 building drawing)

## Slide 4: Additional door
Note)
Canceled.
V5

## Slide 5: Additional equipment
Floor mounted type autoclave
Sink

## Slide 6: Window in inside door
Note) Window size is W300mm and H500mm or more.
Note) In case that clear view is not allowed (Shower room, etc.),
Blind film or other method should be considered.
V5

## Slide 7: Door interlock

## Slide 8: Door lock and Open/Close Sensor
Note1) Door lock had better be electric magnetic type.
Door lock and sensor
Card Reader
Button
Note2)
If access record is needed at not only entering room but also exiting room, Button should be changed to Card Reader.
Note3)
Emergency
Lock cancelation is needed.
V6

## Slide 9: Manometer (Pressure gauge)
Note) Manometer should have proper range.
Manometer
Note)
Added.
Note)
Added.
V6

## Slide 10: Camera
Note) Camera should have proper view range and resolution.
Camera
Monitor
Note)
Moved.
Note)
Moved.
V5

## Slide 11: Air Tightness
Note) Door’s air tightness
should be 6m3/m2hPa or less.
Note)
Revised.
V5

## Slide 12: Heat resistance
Note) Up to 100 degC

## Slide 13: Chemical resistance
Note) For alcohol liquid and Formalin gas.
Note) In case of floor, additionally Sodium hypochlorite and Peracetic acid.
V5

## Slide 14: Illumination
Note) Timer operation is needed for animal.
200lx or more
500lx or more
300lx or more
Note) Use LED light.

## Slide 15: Back up power source
Facility, Equipment or Device | Source | Start up time [min] | Running time [min]
BSC, ABSC | Emergency generator + UPS | around 0 by UPS | around 10 by UPS (for stopping BSC safely)
Isolator | Emergency generator | around 5 | around 60
Refrigerator, Freezer | Emergency generator | around 5 | around 60
Supply fan, Exhaust fan and other equipment for room pressure control | Emergency generator | around 5 | around 20 (for stopping air conditioning system safely)
Lighting equipment in BSL3 and ABSL3 | Emergency generator | around 5 | around 20 (for escape safely)
Others | None | ‘- | ‘-
Note1) Emergency generator in PIHCMC can be used, if it has enough capacity.
Note2) Door lock, Camera, Intercom and other security item should be back-up, if needed.
V6

## Slide 16: Temperature control
40 or less
degC
22 to 26
degC
20 to 28
degC
80 or less
degC

## Slide 17: Humidity control
80 or less
%RH
50 to 70
%RH
40 to 80
%RH
45 to 55
%RH
Note)
Revised.
V5

## Slide 18: Air change rate
1 times/h or more
20 times/h or more
12 times/h or more
Not specified
Note) 6 times/h or more at stand-by mode in case of

## Slide 19: Pressure control
-20 to -40
Pa
-80 to -100
Pa
-50 to -70
Pa
Around 0
Pa

## Slide 20: Air conditioning system (Reference)
BSC
HF
PCD
AHU
MF
RF
HF
VCD
EF
HF,
EF
HF,
EF
HF,
EF
HF
PCD
Abbr.) RF: Rough Filter, MF: Medium Efficiency Filter, HF: HEPA Filter, PCD: Pressure Control Damper,
VCD: Volume Control Damper, AHU: Air Handling Unit, BPD: By-Pass Damper, EF: Exhaust Fan,
SF: Supply Fan, BSC: Bio Safety Cabinet, AC: Autoclave, RA: Room Air conditioner
BSL3, ABSL3
AC
RF
SF
EF
Autoclave room,
Machine room, Storage
RF
SF
EF
Corridor,
Control room, Cage wash room
RA
BPD
V6

## Slide 21: HEPA filter casing (Reference)
Note1) Casing can be manufactured by spot welding and sealing.
Note2) HEPA filter has gasket on bottom side.
Note3) Access door has gasket. Gasket should have no seam.
Note4) Flange on both side of BV can be manufactured by steal plate.
Note5) Both GV should be closed except in sterilization. Fan and hose installed between both GV to circulate sterilization gas.
HEPA filter
PS
BV
Air flow
Access hole
Access door
Abbr.)
GV: Gate Valve,
BV: Butterfly Valve,
PG: Pressure gauge,
PS: Pressure Sensor
PG
GV

## Slide 22: Drainage water system (Reference)
Abbr.)
HW: Hand Wash,
SP: Shower Pan,
DT: Dunk Tank,
AC: Autoclave,
ST: Storage Tank,
HF: HEPA Filter,
P: Pump,
WTS: Water Treatment System
PDS: Public Drainage System
AC
P
Autoclave
room
HW
SP
Lab room, Anteroom
DT
Water treatment room
WTS
ST
PDS
HF
Note1) DT’s drainage water can be connected directly with PDS.
Note2) If AC is anti-bio-hazard type, its drainage water can be connected directly with PDS.
Note3) AC’s drainage water is very hot and has pressure.
Note4) Drainage system piping under floor had better be installed in trench.
Note5) HW, SP should have trap. Its water seal height should be 75mm or more.
Vent
Cage Wash
room
V6
Sink
PDS
HF

## Slide 23: Control & Monitor system (Reference)
Abbr.) PLC: Programmable Logic Controller, Comp: Compressor
Room Pressure | VCD opening ratio | System On/Off | Buzzer On/Off
Room Temperature | PCD opening ratio | Buzzer Reset | AHU Fan On/Off
Room Humidity | BPD opening ratio | Door Open/Close | AHU Comp. On/Off
Filter Pressure | Humidifier control | AHU Alarm | EF On/Off
… | … | BSC On/Off | Door Lock On/Off
… | … | … | …
Analog Input | Analog Output | Digital Input | Digital Output
PLC |  |  | 
Graphic Panel (Local) |  | Ethernet (Remote) | 

## Slide 24: CO2 gas supply
CO2 gas cylinders
Note) CO2 gas piping should include Manifold, Pressure reducing valve, Pressure gauge.
CO2 gas piping
V5

## Slide 25: Access flow
V6
For person normally
For person in case of high risk
(1 way)
For large equipment

## Slide 26: Intercom
Note) Extension phone is better, if possible.
Intercom
V6

## Slide 27: Waste water treatment system (Reference)
Abbr.)
HW: Hand Wash,
SP: Shower Pan,
DT: Dunk Tank,
AC: Autoclave,
ST: Storage Tank,
HF: HEPA Filter,
P: Pump,
WTS: Water Treatment System
PDS: Public Drainage System
AC
P
Autoclave
room
HW
SP
Lab room, Anteroom
DT
Water treatment room
WTS
ST
PDS
HF
Note1) DT’s drainage water can be connected directly with PDS.
Note2) If AC is anti-bio-hazard type, its drainage water can be connected directly with PDS.
Note3) AC’s drainage water is very hot and has pressure.
Note4) Drainage system piping under floor had better be installed in trench.
Note5) HW, SP should have trap. Its water seal height should be 75mm or more.
Vent
Cage Wash
room
V7
Sink
PDS
HF